**Journal of Exercise Physiologyonline**

**Volume 14 Number 2 April 2011**



**JEPonline**

**Editor-in-Chief**

Tommy Boone, PhD, MBA

**Review Board**

Todd Astorino, PhD

Julien Baker, PhD

Steve Brock, PhD

Lance Dalleck, PhD

Eric Goulet, PhD

Robert Gotshall, PhD

Alexander Hutchison, PhD

M. Knight-Maloney, PhD

Len Kravitz, PhD

James Laskin, PhD

Yit Aun Lim, PhD

Lonnie Lowery, PhD

Derek Marks, PhD

Cristine Mermier, PhD

Robert Robergs, PhD

Chantal Vella, PhD

Dale Wagner, PhD

Frank Wyatt, PhD

Ben Zhou, PhD

Official Research Journal of the American Society of Exercise Physiologists

ISSN 1097-9751

**Comparing a** **Portable and Standard Metabolic Measuring System during Rest and Exercise**

JASON M. WINKLE1,BLANCHE W. EVANS2, PETER DILG3, SARAH GALPAROLI4

1-4Indiana State University,Terre Haute, Indiana, USA

##### ABSTRACT

**Winkle JM, Evans BW, Dilg P, Galparoli S.** Comparing a Portable and Standard Metabolic Measuring System during Rest and Exercise. **JEPonline**2011;14(2):80-87. The usefulness of a portable metabolic measuring system is dependent upon its ability to produce valid and reliable results when compared to standard laboratory equipment such as non-portable metabolic measuring systems. Results from studies comparing the VO2000 with non-portable systems have been inconsistent. The purpose of this study was to compare the metabolic and ventilatory values during exercise using a portable metabolic measuring system (VO2000) with a facemask (FM) and mouthpiece/ nose clip (MP) and a standard laboratory metabolic measuring system, CPX-Ultima (CPX/U). Eighteen volunteers (age = 25.7 ± 7.2 yrs) completed three 20 min exercise sessions (10 min walking/10 min jogging) using the CPX/U, FM, and MP. A counter-balanced design determined treatment order. Metabolic and ventilatory values were recorded at 20 sec intervals and averaged over the last 5 min of walking and jogging. A one-way ANOVA with repeated measures compared the mean values for each variable for the three conditions during walking and jogging. No significant difference (p > 0.05) in HR or VO2 was found due to treatment order. The VO2000 produced similar metabolic and ventilatory values whether using the FM or MP (p > 0.05). However, significant differences (p < 0.01) were found for VE, VO2, VCO2, and RER among the CPX/U, FM and MP during walking and jogging. The VO2000 overestimated VO2 compared to the CPX/U by 2.4 mL.kg-1.min-1 during walking and 6.9 mL.kg-1.min-1 during jogging.

**Key Words:** Oxygen Uptake, Portable Metabolic Measuring System

**INTRODUCTION**

Metabolic measuring systems have been used for decades in exercise physiology laboratories to determine physiological and ventilatory responses during exercise. With advances in technology, computerized metabolic measurement systems have been developed that not only produce valid and reliable results, but also provide sophisticated analyses such as estimates of ventilatory threshold and breath-by-breath sampling for oxygen uptake kinetics. Despite the advantages of computerized metabolic measuring systems, a disadvantage has been the lack of portability and, therefore, an inability to directly assess oxygen uptake (VO2) and energy expenditure in the field.

In recent years portable metabolic measuring systems have been designed that enable researchers to obtain measurements during an activity performed outside of a laboratory environment. One portable system, the MedGraphics VO2000, has been used successfully in the field to obtain energy expenditure values for disabled and non-disabled skiers and snow boarders (13), during orienteering (15), and during racquetball competition (3). The usefulness of a portable metabolic measuring system is dependent upon its ability to produce valid and reliable results when compared to standard laboratory equipment such as non-portable metabolic measuring systems. Researchers have reported similar values for basal metabolic rate (BMR) using the VO2000 and a non-portable system specifically designed to measure BMR (17), and also when BMR was measured with the VO2000 using a low flow pneumotach in series with a non-portable metabolic measuring system, the MedGraphics CPX/D (12). During exercise, researchers who compared the VO2000 and the CPX/D in series found no significant differences between the two systems for ventilation (VE), VO2 (L.min-1), and VCO2 (L.min-1) values while cycling at different workloads (4).

On the other hand, when measured during separate exercise sessions, the VO2000 has been reported to be less accurate and less reliable compared to the CPX/D and Douglas bag, overestimating VE, VO2, and VCO2 values over a range of cycling workloads (5). Significantly lower VO2 (mL.kg-1.min-1) and VCO2 (L.min-1) values have been reported during treadmill exercise when using the VO2000 compared to the CPX/D (8). Anderson (2), however, found the VO2000 to be reliable during submaximal and maximal treadmill exercise, but valid for VO2 only at 3.5 mph and at peak intensities (2). The VO2000 was found to overestimate VO2 (mL.kg-1.min-1) compared to the CPX/D during exercise in sedentary pregnant women (20).

Results from studies comparing the VO2000 with standard, non-portable systems during exercise are not consistent and may indicate methodological problems or possible mechanical and technological errors. The over-estimation of values measured in the field by the VO2000 could result in a misrepresentation of actual energy expenditure and VO2 values, resulting in errors in exercise prescription and physical training programs. More research is needed to determine the ability of the VO2000 and similar portable metabolic measuring systems to accurately assess metabolic demands in the field. The purpose of this study was to compare values from a non-portable metabolic measuring system, the MedGraphics CPX Ultima (CPX/U), and a portable metabolic measuring system, the MedGraphics VO2000, during rest, walking, and jogging.

**METHODS**

**Subjects**

Eighteen apparently healthy volunteers (10 males and 8 females), who met the American College of Sports Medicine criteria as low risk (1), were exercising a minimum of three times each week participated in the study. The research procedures used to collect the data were approved by the University Institutional Review Board. Each participant completed a medical history questionnaire, physical activity readiness questionnaire (PAR-Q), and provided written informed consent prior to testing. All subjects were familiar with treadmill exercise and the laboratory procedures. Physical characteristics (mean ± SD) of the subjects were: age = 25.7 ± 7.2 yrs; height = 173.8 ± 9.5 cm; and body mass = 74.0 ± 16.5 kg.

**Equipment**

The CPX/U is a non-portable metabolic measuring system that uses a zirconia O2 analyzer with a response time of <0.80 msec and accuracy of ±0.03% and a non-dispersive infrared (NDIR) CO2 analyzer with response time of <130 msec and accuracy of ±0.1%. A bi-directional differential pressure preVent™ pneumotach with 39 ml dead space and accuracy of ±3% or 50 ml measures ventilation from expiratory volumes over an exercise range of 20 to 200 L.min-1. The BreezeSuite software is used for data collection and manipulation (9).

The VO2000 is a portable metabolic measuring system that uses a three-point harness system or single belt configuration for data collection in a field setting. The unit with the battery pack, harness, and cables weighs 1.5 kg. The O2 analyzer is a galvanic fuel cell with accuracy of ±0.1% and the CO2 analyzer is a NDIR analyzer with an accuracy of ±0.2%. The preVent™ pneumotach measures ventilation from expiratory flow volumes and the BreezeSuite software package is used for data collection and manipulation (10,11).

A neoprene facemask used during data collection has an aluminum nosepiece that is adjusted at the bridge of the nose, a silicone coupler in front of the mouth for attaching the pneumotach, and two strips of neoprene with Velcro to secure the facemask to the crown and nape of the head. The facemask comes in four sizes including extra small, small, medium, and large. The correct fit allows for approximately 1 inch of neoprene material under the chin (10).

**Procedures**

Each participant completed three test sessions separated by a minimum of 24 hrs. The three test sessions included measurements using the non-portable CPX/U, VO2000 with facemask (FM), and the VO2000 with mouthpiece and nose clip (MP). Each session consisted of 5 min of seated rest, 10 min of walking at1.33 m.s-1, and 10 min of jogging at a self-selected speed (2.46 ± .28 m.s-1) on a Trackmaster TMX 425CP treadmill. The three sessions were assigned using a counter-balanced treatment order.

Prior to beginning each test session, participants were fitted with a Polar telemetry heart rate (HR) monitor and either a facemask or mouthpiece and nose clip. The facemask was fitted per manufacturer’s instructions and the same mask size was used for the CPX/U and FM trials. Gas exchange and pulmonary ventilation values were recorded using the breath x breath mode with 20 sec average values printed during testing. Values for variables assessed during the test sessions are reported as the mean value for the last 2 min of rest and for the last 5 min of walking and jogging. Heart rate (HR) was recorded at the end of each min for rest and exercise, respectively. The mean HR for the 5 min of rest and for the last 5 min of exercise was used in data analyses. Rating of perceived exertion (RPE) was recorded at min 4, 5, 9, and 10 during exercise. The mean RPE for the last 2 min of exercise was used in data analyses.

Prior to each test, the CPX/U and VO2000 were calibrated according to manufacturer’s instructions. The CPX/U analyzers were calibrated using medical grade gases and air volume calibrated using a calibrated 3-liter syringe over 5 flow rates. The VO2000 uses an auto-calibration mode that calibrates the analyzers using room air and proprietary software. Since manual calibration was not an option with the VO2000 a manual check of the analyzers and the air flow volume was performed prior to and after the FM and MP trials. The analyzers were checked using medical grade gases and a 3-liter rebreathing bag attached to the CAL port of the VO2000. Oxygen and CO2 values were similar to the gas tank values and were not significantly different prior to and after trials (%O2, t =.87, p = .40 and %CO2, t =.25, p = .81, respectively). The air flow check was completed using a calibrated 3-liter syringe by designating the total volume that would be completed during 20 sec. The value was accepted if it was within 10% of the factory set default value. No volume check was rejected.

**Statistical Analysis**

Values for all variables are reported as the mean ± SD. Each dependent variable was analyzed by a one-way repeated measures analysis of variance (ANOVA), using an alpha level of p < .05. If the sphericity assumption was violated, the p value was adjusted using the Huyn-Feldt correction to determine significance. If significant differences were found, pairwise comparisons using the Bonferroni adjustment determined which trials were different.

**RESULTS**

**Trial Order**

No significant differences related to order of testing were found for VO2 mL.kg-1.min-1 and HR values during rest, walking, and jogging, indicating that the counter-balanced treatment design did not influence the test results (p > .05).

**Correlation Coefficients for FM and MP**

Significant correlations were found during walking between the FM and MP for VE (r = .87, p = .000), VO2 L.min-1 (r = .83, p =. 000), VCO2 L.min-1 (r = .75, p = .000), and VO2 mL.kg-1.min-1 (r = .53, p = .02). During jogging the correlation coefficients for the FM and the MP were significant for VE (r =.92, p = .000), VO2 L.min-1 (r = .78, p = .000), VCO2 L.min-1 (r = .74, p = .000), and VO2 mL.kg-1.min-1 (r = .80, p = .000).

**Comparison of Three Conditions during Walking and Jogging**

Mean physiological and ventilatory values and significant paired differences for walking and jogging for the three conditions are shown in Tables 1 and 2, respectively. No significant differences were found among the three conditions for HR and RPE values during walking and jogging (p > .05).

**Table 1. Metabolic and ventilatory values for the three trials during walking (Mean ± SD).**

|  |  |  |  |
| --- | --- | --- | --- |
| **Variable** |  **CPX/U** |  **FM** |  **MP** |
| **VE (L.min-1)** |  **22.1 ± 4.5a** |  |  **23.7 ± 6.2b** |  |  **28.2 ± 4.5ab** |   |
| **Fb (breaths.min-1)** |  **21.4 ± 4.2ab** |  |  **26.3 ± 7.3ac** |  |  **32.8 ± 9.2bc** |   |
| **TV (mL.breath)**  | **1053.9 ± 243.6** |  | **1020.1 ± 323.0** |  | **945.6 ± 297.6** |  |
| **VO2 (mL.kg-1.min-1)** |  **12.0 ± 1.1ab** |  |  **14.4 ± 2.4a** |  |  **15.6 ± 2.6b** |   |
| **VO2 (L.min-1)**  |  **0.879 ± .182ab** |  |  **1.074 ± .268a** |  | **1.150 ± .190b** |   |
| **VCO2 (L.min-1)**  |  **0.789 ± .184a** |  |  **0.859 ± .215** |  | **0.933 ± .170a** |  |
| **RER** |  **0.90 ± .07ab** |  |  **0.80 ± .05a** |  |  **0.80 ± .07b** |   |

\*Similar letters denote significant differences between conditions.

***Walking***

Significant differences among the three conditions were found during walking for VO2 mL.kg-1.min-1 (F2,16 = 18.58, p = 000), VO2 L.min-1 (F2,16 = 24.55, p = .000), VCO2 L.min-1 (F2,16 = 8.30, p = .000), RER (F2,16 = 25.24, p = .000), VE (F2,16 = 29.79, p = .000), and Fb (F2,16 = 24.79, p = .000). No significant difference among the three conditions was found for TV (p > .05). VO2 values were significantly lower for the CPX/U trial compared to the FM and MP trials, but significantly higher for RER (p < .05). Significantly lower VCO2 and VE values were found for the CPX/U trial compared to the MP trial (p < .05). In addition, VE and Fb were significantly higher for the MP trial compared to the FM trial (p < .05) (Table 1).

***Jogging***

Significant differences among the three conditions were found VO2 mL.kg-1.min-1 (F2,16 = 29.29, p = .000), VO2 L.min-1 (F2,16 = 8.47, p = .000), VCO2 L.min-1 (F2,16 = 5.98, p = .006), RER (F2,16 = 25.23, p = .000), VE (F2,16 = 14.39, p = .000), and Fb (F2,16 = 5.16, p = .01). The CPX/U values were significantly lower compared to the MP values for VO2 results, VCO2, VE, and Fb, but significantly higher for RER (p <.05). In addition, a significantly higher RER was found for the CPX/U trial compared to the FM and MP trials (p < .05) (Table 2).

**Table 2. Metabolic and ventilatory values for the three trials during jogging (Mean ± SD).**

|  |  |  |  |
| --- | --- | --- | --- |
| **Variable** |  **CPX/U** |  **FM** |  **MP** |
| **VE (L.min-1)** |  **59.1 ± 15.5a** |  |  **63.8 ± 19.2** |  |  **67.9 ± 17.5a** |  |
| **Fb (breaths.min-1)** |  **35.0 ± 6.0a** |  |  **36.4 ± 6.3** |  |  **38.1 ± 7.8a** |  |
| **TV (mL.breath)**  | **1723.7 ± 505.5** |  | **1781.8 ± 582.6** |  | **1870.2 ± 514.3** |  |
| **VO2 (mL.kg-1.min-1)** |  **30.5 ± 4.4ab** |  |  **37.4 ± 8.0a** |  |  **37.7 ± 6.3b** |  |
| **VO2 (L.min-1)**  |  **2.194 ± .572a** |  |  **2.566 ± 1.071** |  |  **2.761 ± .718a** |  |
| **VCO2 (L.min-1)**  |  **2.212 ± .615a** |  |  **2.390 ± .708** |  |  **2.417 ± .661a** |  |
| **RER** | 1. **± .06ab**
 |  |  **0.88 ± .05a** |  |  **0.89 ± .07b** |   |

\*Similar letters denote significant differences between conditions.

**DISCUSSION**

Several researchers have reported reliable values for the VO2000 and FM during rest (17) and during submaximal and maximal treadmill exercise (2). In this study the VO2000 produced similar VO2, VCO2, and RER values during walking and jogging when using either a FM or MP. No studies were found that reported results using the MP. Though VO2, VCO2, and RER values were similar in this study, some significant differences occurred in VE as a result of differences in TV and Fb. These results were not surprising, as previous researchers have shown that breathing apparatus affects breathing patterns. In particular, the use of MP increases VE due to an increase in Fb (16). In this studyFb was slightly higher for MP compared to FM during walking and jogging, while TV was higher during jogging for MP. The differences that occur in breathing patterns due to breathing apparatus may reflect differences in breathing assembly dead space, resistance to air flow during inspiration and expiration, leaks in the FM as well as mechanical or neural factors. As with previous studies that compared FM and MP, the differences in VE due to breathing pattern did not translate into statistically significant differences in VO2 values (6). Even though a majority of the participants found the MP to be uncomfortable and unappealing, they perceived the exercise sessions to be similar in intensity, reporting the same RPE values for FM and MP.

The validity of the VO2000 has been in question. Several studies have indicated that the VO2000 produces valid measurements during rest (12,17) and in series with a CPX/D (4). Other studies have found an overestimation of VE, VO2, VCO2, and RER values when using the VO2000 and FM compared to the CPX/D or other portable units (5,20). This study found the FM overestimated VO2 and RER values compared to the laboratory standard CPX/U, while MP overestimated all values except TV. During walking, the overestimation of VO2 values was 2.4 mL.kg-1.min-1 higher using the FM compared to the CPX/U. Typically, 2 to 3 mL.kg-1.min-1 difference or a difference of approximately 3 to 5% between tests is acceptable due to biological and day-to-day variability (18). Therefore, the VO2000 with FM appears to be within industry standards for use during lower intensity exercise such as walking.

On the other hand, the significant difference in VO2 values between the CPX/U and the VO2000 during jogging was 6.9 mL.kg-1.min-1. This difference is troubling, yet it supports the findings of Anderson (2) who reported valid results for the VO2000 at 3.5 mph, but not for 5 and 6 mph. The mean speed for jogging in this study was 5.5 mph. The differences reported between walking and jogging combined with the finding by Anderson (2) that the VO2000 was valid at peak exercise suggests a non-systematic error in the VO2000.

One area for concern mentioned in previous studies has been the measurement of VE. Errors that occur in VE translate directly into errors in VO2 but in this study no significant difference was found between CPX/U and FM for VE. However, the significantly lower RER values for the VO2000 compared to the CPX/U could indicate a problem with lag time during gas analysis. A difference in gas transport time and the response time of the analyzers as well as other default settings for the computerized systems could result in differences in reported values between measuring systems (7,14,19). The factory default settings on the VO2000 are not adjustable. Future studies should focus attention on comparing the factory settings between units to determine if they present a possible source of error prior to testing.

The importance of calibration and verification of systems is critical when using different metabolic measuring systems. The autocalibration system of the VO2000 has been suggested as a possible source of error (8). In this study a manual check was completed that verified the analyzer readings before and after testing using medical grade gases and checked air flow volume using a calibrated 3-liter syringe. Therefore, differences between measuring systems in this study appear to be related to some of the technical issues described previously.

### **CONCLUSIONS**

The VO2000 was reliable during walking and jogging when using either FM or MP. The small difference in VO2 values between the CPX/U and FM during walking indicate the VO2000 is within industry standards, even though statistically different, and may be considered valid for this type of low intensity exercise. However, the difference in VO2 values between systems during jogging indicates the VO2000 was not valid at the average speeds used in this study. Continued research is needed to better understand and address problems that may be encountered when using portable metabolic measuring systems in the field or when comparing values from multiple metabolic measuring systems.

Address for correspondence: Winkle JM, Ph.D., College of Nursing, Health, and Human Services, Indiana State University, Terre Haute, Indiana, USA 47809. Phone (812) 237-3113; FAX: (812) 237-4338; Email: Jason.Winkle@indstate.edu.

**REFERENCES**

1. American College of Sports Medicine. ***ACSM’s Guidelines for Exercise Testing and Prescription***. 7th edition. Philadelphia PA: Lippincott, Williams & Wilkins; 2006.
2. Anderson CL. Validity and reliability of the Medical Graphics VO2000, Cosmed Fitmate Pro, and HealthTech MedGem during rest and at submaximal and maximal exercise intensities ***[Masters Thesis].*** South Dakota State University; 94 pp, AAT 1438182; 2006.
3. Berg K, Narazaki D, Latin R, Vincent W, Meisinger M, Sjoberg C, et al. Oxygen cost and energy expenditure of racquetball. ***J Sports Med Phys Fitness*** 2007;47:395-400.
4. Byard AD, Dengel DR. Validity of a portable metabolic measurement system. ***Med Sci Sports Exerc*** 2002;34:S149.
5. Crouter SE, Antczak A, Hudak JR, DellaValle DM, Haas JD. Accuracy and reliability of the ParvoMedics TrueOne2400 and MedGraphics VO2000 metabolic systems. ***Eur J Appl Physiol*** 2006;98:139-151.
6. Evans BW, Potteiger JA. Metabolic and ventilatory responses to submaximal and maxima exercise using different breathing assemblies. ***J Sports Med Phys Fitness*** 1995;35:93-98.
7. Gore C, Clark R, Shipp N, Ver Der Ploeg G, Withers R. CPX/D underestimates VO2 in athletes compared with an automated Douglas bag. ***Med Sci Sports Exerc*** 2003;35:1341-1347.
8. Kautza BC, Kastello GM, Sothmann MS. Validation of MedGraphics VO2000 portable metabolic analyzer and a modified pneumotachometer. ***Med Sci Sports Exerc*** 2004;36:S222.
9. ***MedGraphics BreezeSuite CPX Ultima [User Manual].*** St. Paul, MN: Medical Graphics Corporation; 2001.
10. ***MedGraphics [Informational Product Bulletin]: PreVent Mask Selection and Fitting Instructions, P/N: 140330-001, Rev B.*** St. Paul, MN: Medical Graphics Corporation; 2003.
11. ***MedGraphics VO2000 [Operation Guide]: An Introduction to the VO2000 Metabolic Analysis System.*** St. Paul, MN: Medial Graphics Corporation; 2001.
12. Olson TP, Tracey JE, Dengel DR. Validity of a low-flow pneumotach and portable metabolic system for measurement of basal metabolic rate. ***Med Sci Sports Exerc*** 2003;35:S143.
13. Petrofsky J, Magsino R, Zook S, Kao JK, Meyer J, Laymon M. A comparative study of monoskiers, conventional downhill skiers, and snowboarders. ***Palaestra*** 2002;18:20-28.
14. Prieur F, Castells J, Demos C. A methodology to assess the accuracy of a portable metabolic system (VmaxST™). ***Med Sci Sports Exerc*** 2007;35:879-885.
15. Roberts AD, Rattray B, Linaker KL, Montgomery PG, Kusnanik NW. Validation of a field based incremental test for performance assessment in orienteers. ***Med Sci Sports Exerc*** 2003;35:S90.
16. Sackner JD, Nixon AN, Davis B, Atkins N, Sackner MA. Effects of breathing through external dead space on ventilation at rest and during exercise. ***Am Rev Respir Dis*** 1980;122:933-940.
17. Wahrlick V, Anjos LA, Lohman TG, Going SB. Validation of the VO2000 indirect calorimeter for measuring basal metabolic rate. ***Med Sci Sports Exerc*** 2005;37:S436.
18. Wasserman K, Hansen J, Sue D, Casaburi R, Whipp B. ***Principles of Exercise Testing and Prescription.*** Philadelphia, PA: Lippincott Williams & Wilkins; 1999.
19. Wenos DL, Koceja DM. Comparison of the reliability of two respiratory valves during maximal exercise testing. ***Med Sci Sports Exerc*** 1992;24:S160.
20. Yeo S, Ronis DL, Antonakos CL, Roberts K, Hayashi R. Need for population specific validation of a portable metabolic testing system: A case of sedentary pregnant women. ***J Nurs Meas*** 2005;13:207-218.

# Disclaimer

The opinions expressed in **JEPonline** are those of the authors and are not attributable to **JEPonline**, the editorial staff or the ASEP organization.